1. Install Comps and Burden Resistors
2. Install SD Card containing NET.CNF and boot.bin
boot.bin is an older firmware revision, before per-channel polarity was implemented:

https://github.com/jamead/zpsc-fw/blob/88e941eb0aa405683c866b5cf4078a65b37348e8/builds/BOOT.bin.githash-0x405262F
3. Change IP Address in NET.CNF, if needed.
4-Channel PSCs should be 10.69.26.30-33 (lab{1} to lab{3})
2-Channel PSCs should be 10.69.26.34-35 (lab{4}, lab{5})

Using the wrong IP may cause some issues with CSS panels, as the 2-channel units and 4-channel unit have some PVs that behave differently.

4. Connect cables, AC Power, EVR Fiber, DCCT cable (Beige), and Power Converter (Black) cables to the rear of the chassis. Connect Micro-USB Cable to the front panel. Power up, and wait for D8 LED to flash rapidly.

Viewed from the back, the order for the black cables are:
2	4
1	3

5. In a terminal, enter minicom for the PSC, e.g., “minicom -D /dev/ttyUSB2” – case sensitive!

6. Configure EEPROM for the unit (Settings B, C, D, and E). Then use Option A to verify. This allows you to set 2- or 4-channel, high- or medium-stability, fast or slow bandwidth, and bipolar/unipolar type.

7. Once the EEPROM is configured, open a new terminal, and go to the test directory:
“cd ~/NSLS-II_PSCs/NEXT-II/NSLS2_zPSC_Calibration_and_Test”

8. Execute the launcher script: “python3 launcher.py”

Follow the on-screen prompts for calibration only, test only, or both calibration and test.

Ensure the PSC to be calibrated is warmed up – at least 15-20 minutes powered on, prior to testing, to allow for electronics to reach stable temperature.